National Standard of the People’s Republic of China

GB/T18883-2002

Indoor air quality standard

Issue date 2002-11-19

Implementation date 2003-03-01

Issued by China State Bureau of Quality and Technical Supervision

Ministry of Public Health

State Environmental Protection Administration

Contents

Foreword

1 Scope	1

2 Normative reference	1

3 Terms and definitions	2

4 Indoor air quality standard	2

5 Indoor air quality test	3

Annex A (normative) guide of indoor air monitoring techniques	4

Annex B (normative) test method of Benzene in the indoor air (by capillary gas chromatography)	7

Annex C (normative) test method of total volatile organic compound (TVOC) in the indoor air (by thermal desorption / capillary gas chromatography)	9

Annex D (normative) method of total number of bacterial colony measurement in the indoor air	12

Foreword

This standard is published to protect human health, to prevent and control the indoor air pollution.

The Annex A, Annex B, Annex C and Annex D are the normative annexes.

This standard is issued for the first time.

This standard was drafted together by Ministry of Public Health and indoor air quality standard group of China State Bureau of Quality and Technical Supervision.

The drafting units include: Environmental Health and Related Product Safety Institute of China Disease Prevention and Control Center, Environmental Standards Institute of Chinese Academy of Environmental Sciences, Radiation Protection Safety Institute of China Disease Prevention and Control Center, College of Environmental Sciences of Peking University, College of Environmental Science and Engineering of Nankai University, Beijing Labor Protection Institute, School of Architecture of Tsinghua University, Ecological Environment Research Center of Chinese Academy of Sciences and Institute of Environmental Engineering of China Building Materials Academy.

This standard was approved by General Administration of Quality Supervision, Inspection and Quarantine, Ministry of Public Health and China State Bureau of Quality and Technical Supervision on November 19, 2002.

This standard was proposed by General Administration of Quality Supervision, Inspection and Quarantine.

This standard is interpreted by General Administration of Quality Supervision, Inspection and Quarantine and Ministry of Public Health.

Indoor air quality standard

1	Scope

This standard specifies the indoor air quality parameters and test methods.

This standard is applicable to the residential and office buildings, and may also apply to other indoor environment by reference.

2	Normative reference

The following standards contain provisions which, through reference in this text, constitute provisions of this standard. For dated reference, subsequent amendments to (excluding correction to), or revisions of, any of these publications do not apply. However, the parties to agreements based on this standard are encouraged to investigate the possibility of applying the most recent editions of the standards. For undated references, the latest edition of the normative document referred to applies.

| GB 9801 | Air quality-Determination of carbon monoxide - Non-dispersive infrared spectrometry |
| --- | --- |
| GB 11737 | Standard method for examination of benzene, toluene and xylene in air of residential areas - Gas chromatography |
| GB 12372 | Standard method for examination of nitrogen dioxide in air of residential areas - Modified Saltzman method |
| GB/T 14582 | Standard method for radon measurement in environmental air |
| GB/T 14668 | Air quality - Determination of ammonia - Nesslers reagent colorimetric method |
| GB/T 14669 | Air quality - Determination of ammonia - Ion selective electrode method |
| GB 14677 | Air quality - Determination of toluene, dimethyl benzene and styrene - Gas chromatography |
| GB/T 14679 | Air quality - Determination of ammonia - Sodium salicylate - sodium hypochlorite spectrophotometric method |
| GB/T 15262 | Ambient air - Determination of sulfur dioxide - Formaldehyde absorbing - pararosaniline spectrophotometry |
| GB/T 15435 | Ambient air - Determination of nitrogen dioxide - Saltzman method |
| GB/T 15438 | Ambient air - Determination of ozone - Ultraviolet photometric method |
| GB/T 15439 | Ambient air - Determination of Benzo [a] pyrene in ambient air - High performance liquid chromatography |
| GB/T 15516 | Air quality - Determination of formaldehyde - Acetylacetone spectrophotometric method |
| GB/T 16128 | Standard health inspection method of sulfur dioxide in the air of residential area - formaldehyde solution absorption - pararosaniline hydrochloride spectrophotometry |
| GB/T 16129 | Standard health inspection method of formaldehyde in the air of residential area - spectrophotometry |
| GB/T 16147 | Measurement method of radon concentration in the air - scintillation flask |
| GB/T 17095 | Inhalable particle health standard in the indoor air |
| GB/T 18204.13 | Determination method of indoor temperature in public places |
| GB/T 18204.14 | Determination method of indoor relative humidity in public places |
| GB/T 18204.15 | Determination method of indoor air velocity in public places |
| GB/T 18204.18 | Determination method of indoor fresh air requirement in public places - tracer gas technique |
| GB/T 18204.23 | Test method of carbon monoxide in public places |
| GB/T 18204.24 | Test method of carbon dioxide in public places |
| GB/T 18204.25 | Test method of ammonia in public places |
| GB/T 18204.26 | Determination method of formaldehyde in public places |
| GB/T 18204.27 | Test method of ozone in public places |

3 	Terms and definitions

3.1	Indoor air quality parameter

The physical, chemical, biological and radiological parameters in the air related to human health.

3.2 	Particulates with diameter 10μm or less, PM10)

Particles suspending in the air whose aerodynamic equivalent diameter is less than or equal to 10μm.

3.3 	Total volatile organic compound (TVOC)

TVOC refers to the volatile organic compound whose retention time is between n-hexane and n-hexadecane in the analysis of non-polar column (polarity index of less than 10) sampled with Tenax GC or Tenax TA.

3.4 	Normal state

It refers to dry matter state of temperature of 273K and pressure of 101.325kPa.

4 	Indoor air quality standard

4.1 	Indoor air should be non-toxic, harmless and no abnormal smell.

4.2 	Indoor air quality standard is shown in Table 1.

Table 1 Indoor Air Quality Standard

| NO. | Parameters | Parameters | Unit | Standard | Remarks |
| --- | --- | --- | --- | --- | --- |
| 1 | Physical | Temperature | °C | 22~28 | Summer, air conditioning |
|  |  |  |  | 16~24 | Winter, heating |
| 2 |  | Relative humidity | % | 40~80 | Summer, air conditioning |
|  |  |  |  | 30~60 | Winter, heating |
| 3 |  | Air flow rate | m/s | 0.3 | Summer, air conditioning |
|  |  |  |  | 0.2 | Winter, heating |
| 4 |  | Fresh air | m3/(h.person) | 30a |  |
| 5 | Chemical | sulfur dioxide SO2 | mg/m3 | 0.50 | Average per hour |
| 6 |  | nitrogen dioxide NO2 | mg/m3 | 0.24 | Average per hour |
| 7 |  | carbon monoxide CO | mg/m3 | 10 | Average per hour |
| 8 |  | carbon dioxide CO2 | % | 0.10 | Average per day |
| 9 |  | ammonia NH3 | mg/m3 | 0.20 | Average per hour |
| 10 |  | ozone O3 | mg/m3 | 0.16 | Average per hour |
| 11 |  | formaldehyde HCHO | mg/m3 | 0.10 | Average per hour |
| 12 |  | benzene C6H6 | mg/m3 | 0.11 | Average per hour |
| 13 |  | toluene C7H8 | mg/m3 | 0.20 | Average per hour |
| 14 |  | Xylene C8H10 | mg/m3 | 0.20 | Average per hour |
| 15 |  | benzo [a] pyrene B(a)P | ng/m3 | 1.0 | Average per day |
| 16 |  | inhalable particles PM10 | mg/m3 | 0.15 | Average per day |
| 17 |  | Total Volatile Organic Compounds TVOC | mg/m3 | 0.60 | Average per 8 h |
| 18 | Biological | total number of bacteria colony | cfu/m3 | 2500 | Instrument dependent b |
| 19 | Radiological | radon 222Rn | Bq/m3 | 400 | Year average (action level c) |
| a refresh air requirements should be equal to or greater than standard value, other parameters except temperature and relative humidity requirements should be equal to or less than standard value; b see Annex D; c it is recommended to take measurements to reduce the indoor radon concentration when it reaches this level. |  |  |  |  |  |

5 	Indoor air quality test

5.1 	See Annex A for guide of indoor air monitoring techniques.

5.2 	See Annex B for test method of Benzene in the indoor air (by capillary gas chromatography).

5.3 	See Annex C for test method of total volatile organic compound (TVOC) in the indoor air (by thermal desorption / capillary gas chromatography).

5.4 	See Annex D for method of total number of bacterial colony measurement in the indoor air.

Annex A

(Normative)

Guide of indoor air monitoring techniques

A.1 	Scope

This guide provides the air monitoring requirements of sampling points, sampling time and frequency, sampling methods, measurement methods of instruments and various indoor air quality parameters, quality assurance measures and test results and evaluations.

A.2 	Sampling points

A.2.1	Number of sampling points: The number of sampling points should be determined according to the size of monitor room and scene in order to correctly reflect the level of indoor air pollutants. Generally, 1~3 points should be located in the room less than 50m2; 3~5 points in the room 50~100m2; 5 points in the room larger than 100m2. The points should be distributed in the diagonal or in the plum blossom type pattern.

A.2.2	Sampling points should be away from the vents and more than 0.5m away from the walls.

A.2.3	Sampling points height: Should be the same height as human breathing zone, relative height between 0.5m and 1.5m.

A.3 	Sampling time and frequency

The sampling time should be more than 3 months for annual concentration, 18h for daily concentration, 6 hours for average concentration per 8h, 45min for average concentration per hour. The sampling time should cover the worst ventilation period.

A.4 	Sampling methods and instruments

Choose appropriate sampling methods and instruments according to the state of existence of pollutants in the indoor air. The sampler noise should be less than 50dB (A). The specific sampling methods should be carried out according to the test methods of each pollutant.

A.4.1	Screen sampling: Close the doors and windows for 12 hours. Close the doors and windows when sampling. Sample at least 45min.

A.4.2	Accumulation sampling: When the screen sampling cannot meet the requirements of this standard, accumulation sampling should be used (annual average, daily average and 8h average).

A.5	Quality assurance measures

A.5.1	Seal check: Check the seal tightness of the sampling system before sampling and make sure there is no leak.

A.5.2	Flow calibration: Sampling system should keep constant flow. Use the first level soap film flowmeter to calibrate the intake flow of the sampling system before and after sampling. The deviation error should be within 5%.

Sampler flow calibration: Use the first level soap film flowmeter to calibrate the sampling system with 5 points and draw the standard flow curve when in its normal working state. Record the atmospheric pressure and temperature when calibrating.

A.5.3	blank check: Reserve two sampling tubes for blank check. If the blank check result is over the normal range, then stop the test and change another batch of samples.

A.5.4	Test and calibrate the instruments according to the instrument instructions before use.

A.5.5	Translate the sampling volume to the standard state volume when calculating the concentration using the formula below:

Wherein	V0	converted standard state volume of sample, L;

V	volume of sample, L;

T0	absolute temperature in the standard state, 273K;

T	sum of the sampling temperature (t) and absolute temperature in the standard state, (t+273) K;

P0	atmospheric pressure in the standard state, 101.3kPa;

P	atmospheric pressure of the sampling point, kPa.

A.5.6	Do the parallel sampling each time and the relative difference between the test results and the average value should be within 20%.

A.6	Test methods

Test methods of the parameters of the indoor air are shown in Table A.1.

Table A.1   Test methods of the indoor air parameters

| NO. | Pollutant | Test method | Source |
| --- | --- | --- | --- |
| 1 | sulfur dioxide SO2 | (1) Formaldehyde solution absorption - hydrochloric pararosaniline spectrophotometry | (1) GB/T 16128 (2) GB/T 15262 |
| 2 | nitrogen dioxide NO2 | (1) improved Saltzaman method | (1) GB/ 12372 (2) GB/T 15435 |
| 3 | carbon monoxide CO | (1) non-dispersive infrared method (2) non-dispersive infrared gas analysis, gas chromatography, mercury displacement method | (1) GB 9801 (2) GB/T 18204.23 |
| 4 | carbon dioxide CO2 | (1) non-dispersive infrared gas analysis method (2) gas chromatography (3) capacity titration method | GB/T 18204.24 |
| 5 | ammonia NH3 | (1) indophenol blue spectrophotometric method, Nessler reagent spectrophotometry (2) ion selective electrode method (3) sodium hypochlorite - salicylic acid spectrophotometry | (1) GB/T 18204.25 (2) GB/T 14669 (3) GB/T 14679 |
| 6 | ozone O3 | (1) ultraviolet spectrophotometry (2) indigo carmine spectrophotometry | (1) GB/T 15438 (2) GB/T 18204.27 GB/T 15437 |
| 7 | formaldehyde HCHO | (1) AHMT spectrophotometry (2) phenol reagent spectrophotometry, gas chromatography (3) acetylacetone spectrophotometry | (1) GB/T 16129 (2) GB/T 18204.26 (3) GB/T 15516 |
| 8 | benzene C6H6 | gas chromatography | (1) Annex B (2) GB 11737 |
| 9 | toluene C7H8 xylene C8H10 | gas chromatography | (1) GB 11737 (2) GB 14677 |
| 10 | benzo [a] pyrene B(a)P | high pressure liquid chromatography | GB/T 15439 |
| 11 | inhalable particles PM10 | percussive - weighing method | GB/T 17095 |
| 12 | total volatile organic compounds TVOC | gas chromatography | Annex C |
| 13 | total number of bacteria | impacting method | Annex D |
| 14 | temperature | (1) liquid-in-glass thermometer method (2) digital thermometer method | GB/T18204.13 |
| 15 | relative humidity | (1) ventilation psychrometer method (2) lithium chloride hygrometers method (3) capacitive digital hygrometers method | GB/T18204.14 |
| 16 | air flow rate | (1) hot bulb electric anemometer method (2) digital anemometer method | GB/T18204.15 |
| 17 | fresh air | tracer gas method | GB/T18204.18 |
| 18 | radon222Rn | (1) radon concentration measurement method with scintillation flask (2) track etching method (3) double-filter method (4) activated charcoal method | (1) GB/T 16147 (2) GB/T 14582 |

A.7	Record

Record the scene conditions, various pollution sources, sampling date, time, location, number of sampling points, distribution pattern, atmospheric pressure, temperature, relative humidity, air flow rate and tester signature, etc. and report to the laboratory together with the sample.

Detailed records should be checked including check date, lab, instruments, number, analysis method, test principle, experimental conditions, raw data, tester, and checker.

A.8	Test results and evaluation

The test results are compliance with this standard if the chemical, biological and radiological indicators in the mean value meet each requirement. If one or more of the test results does not meet the requirements of its requirements, then it does not meet the requirements of this standard.

For the parameters in annual average, daily average and 8h average, screen sampling could be carried out. If the test result meets its requirements, it meets the requirements of this standard. If the test result does not meet its requirements, accumulation sampling test should be carried out and tested according to the procedure of annual average, daily average and 8h average parameters.

Annex B

(Normative)

Determination of benzene concentration in the air

(Capillary gas chromatography)

B.1 	Method summary

B.1.1 	Related standard

This method is mainly based on the GB 11737-89 standard test methods of benzene, toluene and xylene in the air of residential areas - gas chromatography.

B.1.2 	Principle: Collect the benzene in the air with activated carbon tube and extract with carbon disulfide. Analyze it with flame ionization detector of gas chromatograph to keep the qualitative of time and quantitative of peak.

B.1.3 	Interference and elimination: Too much water vapor or mist in the air could affect the breakthrough capacity and sampling efficiency of the activated carbon. Humidity should be under 90% in the air for the requirements of sampling efficiency. Other interference in the air could be eliminated by appropriate separation conditions using gas chromatography techniques.

B.2	Scope

B.2.1 	Range: The sampling volume is 20L, extract with 1ml carbon disulfide, sample injection of 1µl, range of 0.05~10 mg/m3.

B.2.2 	Applicable areas: It is applicable for the determination of benzene concentration in the air of indoor or residential areas.

B.3 	Reagents and materials

B.3.1 	benzene: chromatographically pure.

B.3.2	carbon disulfide: analytically pure, purification, to ensure no mixed peaks in chromatographic analysis.

B.3.3 	coconut shell activated carbon: 20~40, for load of activated carbon sampling tube.

B.3.4 	pure nitrogen: 99.999%.

B.4 	Instruments and equipment

B.4.1 	activated carbon sampling tube: length of 150mm, inner diameter of 3.5~4.0mm, outer diameter of 6mm. Put 100mg coconut shell activated carbon in the tube and fix both ends with small amount of glass wool. Blow the tube for 5~10 min in the pure nitrogen of 300~350°C and seal the both ends of the tube with plastic cap. Place the tube in the dryer for 5d. If the tube is sealed with heat, it may be stable for three months.

B.4.2 	air sampler: flow range of 0.2~1L/min, stable. Calibrate the flow of the sampling system with soap film flowmeter before and after sampling. The error should be less than 5%.

B.4.3 	injector: 1ml. Volume scale error should be corrected.

B.4.4 	micro injector: 1µl, 10µl. Volume scale error should be corrected.

B.4.5 	tube with scale and cap: 2ml.

B.4.6 	gas chromatograph: with flame ionization detector.

B.4.7 	columns: 0.53mm×30mm large-caliber non-polar silica capillary column.

B.5 	Sample and preservation

Open the tube in the scene. The size of holes of both ends should be at least 2mm. Connect with the sampler vertically and extract 20L air in the speed of 0.5L/min. After sampling, plug the plastic cap into both ends of the tube and record the temperature and atmospheric pressure. The sample can be saved for 5d.

B.6 	Analysis procedure

B.6.1 	chromatography analysis condition: The chromatography analysis condition varies with the different experimental conditions. Optimal chromatography analysis condition should be determined according to the model and performance of the gas chromatograph.

B.6.2 	standard curve and the factor: Draw the standard curve and calculate the factor under the same condition with the sample analysis.

Draw the standard curve with standard solution: Add a small amount of carbon disulfide in the flask of 5.0ml and certain benzene (1µL of benzene weights 0.8787g in 20 °C) with 1µL micro injector. Add carbon disulfide to the mark to formulate stock solution of certain concentration. Dilute the stock solution separately to form the standard solution of benzene concentration of 2.0, 5.0, 10.0 and 50.0 µg/ml. Take 1µL standard solution to test the preservation time and peak. Repeat for 3 times for each concentration and take the average value of the peak. Draw the standard curve with 1µL benzene content (µg/ml) as abscissa (µg), and average peak as ordinate (mm). Calculate the slope of the regression line and use the inverse of the slope Bs[µg/mm] as the factor of the sample.

B.6.3 	sample analysis: Put the activated carbon into the tube with scale, add 1.0ml carbon disulfide, plug the tube and place for 1h. Shake at times. Take 1µl to keep the qualitative of time and quantitative of peak. Analyze each sample for 3 times and calculate the average of the peak values. At the same time, do the same procedure for an unsampled activated carbon, and measure the average peak value of the blank tube (mm).

B.7 	Result calculation

B.7.1 Convert the sample volume to the normal state volume with formula (1)

┄┄┄┄┄┄┄┄┄┄┄┄┄┄┄┄(1)

Wherein	V0	converted standard state volume of sample, L;

V	volume of sample, L;

T0	absolute temperature in the standard state, 273K;

T	sum of the sampling temperature (t) and absolute temperature in the standard state, (t+273) K;

P0	atmospheric pressure in the standard state, 101.3kPa;

P	atmospheric pressure of the sampling point, kPa.

B.7.2 Calculate the benzene concentration with formula (2):

┄┄┄┄┄┄┄┄┄┄┄┄┄┄┄┄(2)

Wherein	c	benzene, toluene, xylene concentration in the air, mg/m3;

h	average peak value, mm;

h'	average peak value of the blank tube, mm;

Bs	calculated factor from 6.2, µg/mm;

Es	extraction efficiency of carbon disulfide determined by experiment;

V0	converted standard state volume of sample, L.

B.8 	Method characteristics

B.8.1 	detection lower limit: sample volume of 20L, extract with 1ml carbon disulfide, sample injection of 1µl, detection lower limit of 0.05 mg/m3.

B.8.2 	linear range: 106.

B.8.3 	precision: benzene concentration of 8.78 and 21.9µg/ml, relative standard deviation of 7% and 5%.

B.8.4	accuracy: recover rate of 95%, 94% and 91% for 0.5, 21.1 and 200µg benzene respectively.

Annex C

(Normative)

Test method of total volatile organic compound (TVOC) in the indoor air

(Thermal desorption / capillary gas chromatography)

C.1 Method summary

C.1.1 Related standard

ISO 16017-1 “Indoor，ambient and workplace air—Sampling and analysis of volatile organic compounds by sorbent tube/thermal desorption/capillary gas chromatography—part 1：pumped sampling”

C.1.2 Principle

Choose a suitable adsorbent (Tenax GC or Tenax TA). Collect a certain volume of air sample with adsorption tube. The volatile organic compounds in the air flow are retained in the adsorption tube. Heat the adsorption tube and desorb the volatile organic compounds. Put the sample together with the inert carrier gas into the capillary gas chromatograph. Keep the qualitative of time and quantitative of peak.

C.1.3 Interference and elimination

Process and activate the sampling tube and absorbent before sampling to minimize the interference; choose appropriate columns and analysis conditions to separate a variety of volatile organic compounds and solve the problems of interference.

C.2 Scope

C.2.1	Measure range: this method is suitable for the determination of VOCS in the air in the concentration range of 0.5mg/m3~100mg/m3.

C.2.2 	Applicable areas: this method is applicable for the air in indoor, environmental and worksite areas, or the release of small or large cabin material test.

C.3 	Reagents and materials

The reagents used in the analysis process should be chromatographically pure; analytical pure level reagents should be purified to ensure no mixed peaks in the chromatographic analysis.

C.3.1 VOCS: Use VOCS as the reference reagent to calibrate the concentration to form the standard solution and gas. Use liquid external standard method or gas external standard method to inject quantitative solution.

C.3.2 dilution solvent: dilution solvent used in liquid external standard method should be chromatographically pure and separated with the test compound in the chromatographic elution profile.

C.3.3 adsorbent: the diameter of the adsorbent is 0.18~0.25mm (60~80). The adsorbent should be heated with an inert gas stream and activation treated overnight at its highest use temperature. The adsorbent should be cool to room temperature, storage and add to tube in the clean air in order to prevent secondary pollution. The desorption temperature should be below the activation temperature. The adsorbent tube packed by the manufacturers should also need activation treatment before use.

C.3.4 pure nitrogen: 99.999%.

C.4 	Instruments and equipment

C.4.1 adsorption tube: stainless steel tube with polished inner wall, outer diameter of 6.3mm, inner diameter of 5mm, length of 90mm. There is a mark at sampling end of the adsorption tube. Adsorption tube could filled with one or more adsorbents and place the adsorption layer in the heating zone of the desorption instrument. Adsorption tube could fill 200~1000mg adsorbents according to its density. The two ends of the tube should be blocked with stainless steel mesh or glass fiber. If there are more than one adsorbent, the adsorbents should be arranged by its adsorption capacity. The adsorbent should be separated with glass fiber wool. The adsorbent with the least adsorption capacity should be located at the end of sampling end of the adsorption tube.

C.4.2 injector: 10mL liquid injector of 0.1mL precision; 10mL gas injector of 0.1mL precision; 1mL liquid injector of 0.01mL precision.

C.4.3 sampling pump: single constant current air pump, flow range of 0.02~0.5L/min, stable flow. Calibrate the flow of the sampling system with soap film flowmeter before and after sampling. The flow error should be less than 5%.

C.4.4 gas chromatograph: with flame ionization detector, mass detector or other suitable detector.

column: non- polar (polarity index less than 10) fused silica capillary column.

C.4.5 thermal desorption instrument: secondary thermal desorption capable, and bring the desorption gas in to the gas chromatograph with the inert gas. Desorption temperature, time and gas flow rate should be adjustable. The cold trap can concentrate the desorbed samples.

C.4.6 injection device for the liquid external standard: conventional gas chromatographic inlet, online or independent assembly, inlet carrier gas connection preserved, and the lower end of the inlet connected with the adsorption tube.

C.5 	Sampling and preservation

Connect the adsorption tube with sampling pump with plastic or silicone rubber tube. Mount the sampling tube vertically in the breaching zone. Choose the appropriate sampling position for the fixed position sampling. Open the sampling pump and adjust the flow rate to ensure to obtain the desired sampling volume (1~10L) in the appropriate time. If the weight of the total samples is over 1mg, the sampling volume should be reduced respectively. Record the sampling start time and finish time, sampling flow rate, temperature and atmospheric pressure.

Remove the tube after sampling, seal both ends of the tube or put it into the sealed metal or glass containers. The samples can be stored for 14 days.

C.6 	Analysis procedure

C.6.1 Desorption and concentration of the sample

Mount the adsorption tube on the thermal desorption instrument, heat the instrument to make the organic vapor desorb from the adsorbent and be brought into the cold trap by the carrier gas flow and pre-concentrate. The flow direction of the carrier should be the opposite to the direction of sampling. Do the fast desorption in low rate flow and transmit into the capillary gas chromatograph. The temperature of the transmission line should be high enough to prevent condense of the determinand. Desorption conditions (see Table C.1).

Table C.1    Desorption conditions

| Desorption temperature | 250~325°C |
| --- | --- |
| Desorption time | 5~15min |
| Desorbed gas flow rate | 30~50ml/min |
| Refrigeration temperature of cold trap | +20~-180°C |
| Heating temperature of cold trap | 250~350°C |
| Adsorbent in the cold trap | Same as adsorption tube if used, 40~100mg |
| Carrier gas | Helium or high purity nitrogen |
| Split ratio | The split ratio between sampling tube and secondary cold trap and secondary cold trap and analysis columns should be determined by the concentration in the air. |

C.6.2 Chromatographic conditions analysis

Choose the 50m×0.22mm quartz column of 1～5μm thickness. The stationary phase can be polydimethylsiloxane or 7% of the cyano-propane, 7% phenyl and 86% methylsiloxane. The column heat operation should be programed and keep the initial temperature of 50 °C for 10 min and heat to 250 °C with the speed of 5 °C/min.

C.6.3 Standard curve

Gas external standard method: extract 100μg/m3 standard gas of 100ml, 200ml, 400ml, 1L, 2L, 4L and 10L accurately with pump. Prepare the standard with adsorption tube.

Liquid external standard method: use the sample injection device in 4.6 to take 1~5μl standard solution of 100μg/ml and 10μg/ml into the adsorption tube and let the inert gas of 100ml/min go through the adsorption tube. Take off the adsorption tube after 5min and seal the tube to prepare the standard series.

Analyze the adsorption tube standard series with thermal desorption gas chromatography analysis method and draw the standard curve with logarithm of the peak area after blank removing as ordinate and the logarithm of the mass of the analyte as abscissa.

C.6.4 Sample analysis

Analyze each sample adsorption tube according to the drawing procedure of the standard curve (the same desorption and concentrate conditions and chromatographic analysis conditions). Keep the qualitative of time and quantitative of peak area.

C.7 	Result calculation

C.7.1 Convert the sample volume to the normal state volume with formula (1)

┄┄┄┄┄┄┄┄┄┄┄┄┄┄┄┄(1)

Wherein	V0	converted standard state volume of sample, L;

V	volume of sample, L;

T0	absolute temperature in the standard state, 273K;

T	sum of the sampling temperature (t) and absolute temperature in the standard state, (t+273) K;

P0	atmospheric pressure in the standard state, 101.3kPa;

P	atmospheric pressure of the sampling point, kPa.

C.7.2 TVOC calculation

  - All compounds whose preserve time is between n-hexane and n-hexadecane should be analyzed.

  - Calculate TVOC including all compounds between n-hexane and n-hexadecane in the chromatogram.

  - Try to quantitate as many as possible VOCS according to the single calibration curve. At least 10 peak should be quantitated. List the name and concentration of the compounds including TVOC.\

  - Calculate the concentration Sid of the identified and quantitative VOCs.

  - Calculate the concentration Sun of the unidentified VOCs with toluene response coefficient.

  - The sum of Sid and Sun is the concentration value of TVOC.

  - If the VOC value is out of the range defined in (2), add the information to the TVOC.

C.7.3 The concentration of the analytes in the sample air should be calculated with formula (2)

┄┄┄┄┄┄┄┄┄┄┄┄┄┄┄┄(2)

Wherein	c	concentration of the analytes in the sample air, μg /m3;

F	mass of the analytes in the sample tube, μg ;

B	mass of the analytes in the blank tube, μg;

V0	converted standard state volume of sample, L.

C.8 Method characteristics

C.8.1 detection lower limit: detection lower limit of 0.5 μg/m3 when the sampling volume is 10L.

C.8.2 linear range: 106.

C.8.3 precision: add 10µg mixed standard solution in the adsorption tube, the relative standard deviation of Tenax TA should be within 0.4%~2.8%.

C.8.4 accuracy: in the conditions of 20°C and relative humidity of 50%, and n-hexane of 10mg/m3 in the adsorption tube, the total uncertainty of Tenax TA and Tenax GR (average of 5 tests) should be 8.9%.

Annex D

(Normative)

Method of total number of bacterial colony measurement in the indoor air

D.1 	Scope

This method is applicable to the measurement of total number of bacterial colony in the indoor air.

D.2 	Definition

Impacting method is a sampling test method to calculate the total number of bacterial colony per cubic meter of air. It is sampled with percussive microbial sampler, produce high speed current by pumping through slit or hole and make the particles with bacteria suspended in the air strike on the nutrient agar plates. It is calculated after 48 hours of incubation in 37°C.

D.3 Instruments and equipment

D.3.1 Autoclaves

D.3.2 Dry heat sterilization

D.3.3 Incubator

D.3.4 Refrigerator

D.3.5 Plate

D.3.6 Equipment for medium: graduated cylinder, Erlenmeyer flask, pH meter or precise pH paper, etc.

D.3.7 Percussive microbial sampler

Basic sampler requirements:

(1) Capture rate of 95% for the bacteria in the air.

(2) Easy to use, portable, stable and easy to disinfect.

D.4 Nutrient agar medium

D.4.1Ingredients:

Peptone    		20g

Beef extract 	 	3g

Sodium chloride   	5g

Agar	      		15~20g

Distilled water    	1000ml

D.4.2 Steps: mix the ingredients above, heat to dissolve, calibrate the pH value to 7.4, filter and dispense, 121°C, 20min of autoclaving. Impact method refers to the sampler instructions to prepare the nutrient agar plates.

D.5 	Procedures

D.5.1 Choose representative room and location as sampling points. Disinfect the sampler and sample according to the instructions.

D.5.2 After sampling put the nutrient agar plate in the incubator of 36±1°C for 48h. Calculate the number of colonies and convert to the number of colonies per cubic meter of air according to the sampler flow rate and sampling time. Report the results in cfu/m3.